## Slide 1: Fluorescence Microscopy Kurt Thorn NIC

## Slide 2: Why fluorescence?
High contrast
  Signal against dark background
Highly specific, multi-color labeling
  GFP etc.
  Antibodies
Live imaging
  GFP etc.
Quantitative
Sensors for [Ca], pH, …

## Slide 3: What is fluorescence?
Emission light is longer wavelength (lower energy) than excitation light

## Slide 4
Jablonski diagram
(Molecular energy diagram)
S0
S1
S2
em
fluorescence
exc
Lifetime
 ~ 1–4 ns
Singlet states
Spin S=0

## Slide 5
Jablonski diagram
(Molecular energy diagram)
S0
S1
S2
em
T1
T2
fluorescence
exc
Lifetime
 ~ 1–4 ns
Singlet states
Spin S=0
Triplet states
Spin S=1

## Slide 6
Jablonski diagram
(Molecular energy diagram)
Singlet states
Spin S=0
Triplet states
Spin S=1
S0
S1
S2
phospho
exc
T1
T2
phosphorescence
Inter-
systems
crossing
Lifetime
µs–ms

## Slide 7
Total spin S=0
Singlet state
Singlet and Triplet States
Highest filled orbital
Orbital states of each electron
Lowest unfilled orbital
exc
Total spin S=1
Triplet state
Spin flip
Spin flips are “dipole forbidden”  unlikely  long triplet lifetime

## Slide 8: Fluorescence Spectra
Stokes shift
Alexa 488
Excitation
maximum
Emission
maximum

## Slide 9: The Epifluorescence Microscope

## Slide 11
Bandpass
Longpass

## Slide 12
OD
= -log10(%T)
%T

## Slide 13: Interference filters
Filter makers : Chroma
Semrock
Omega
Interference
 Wavelength-
Dependent
transmission
& reflection
 filter
n2
n1
n3
n4
n6
n7
n5
n8

## Slide 14: Filter spectrum is angle dependent
Semrock
Generally bad

## Slide 15: Filter spectrum is angle dependent
Semrock
But can be good – Semrock VersaChrome

## Slide 16: Blocking
Interference filters have finite stop bands
Semrock 697/75

Transmission
To block unwanted transmission from UV to IR, filter makers add
absorption glass to the filter.
Unblocked bandpass interference filter
Pass
band
Stop
band
Stop
band
Often excitation filters are blocked,
but emission filters unblocked.
Red autofluorescence or room light
may get through your blue emission filter
Exc.
pass
band
UV absorber
Interference filter
IR absorber
100%

## Slide 17: Filter cube nomenclature
Chroma labels filters as center wavelength / passband (e.g. D350/50x)
Dichroics are labeled by cut-on wavelength (e.g. 505DCLP)
Nikon filters use a letter to specify illumination wavelength – e.g. UV, B, G, R
Letters afterward specify emission profile – e.g. UV-2A vs UV-2E/C

## Slide 18: Exploding the filter cube
Emission Filter
Dichroic Mirror
Excitation Filter (off slide)

## Slide 19
Filter schemes
Single wavelength sets
  Most efficient
  Best separation
  Very slow to change 
Multi-band filters
Multi-band everything
  See all colors at once
  For color cameras
  Bad crosstalk
“Pinkel” scheme
  Multi-band dichroic
  Multi-band emitter
  Single-l exciters
  Excitaton filter wheel
  Separate image at each wavelength
  Better separation
“Sedat” scheme
  Multi-band dichroic
  single-band emitters
  Single-l exciters
  Two filter wheels
  Even better separation
Chroma triple Pinkel set
Semrock quad Sedat set
Wavelength
Transmission

## Slide 20: Light Sources
Arc Lamps
  Hg and Xe
  Metal Halide
LEDs
Plasma
Lasers
  Generally only for collimated illumination (Confocal, TIRF)

## Slide 21: Arc Lamp Spectra
Hg: Brighter if your excitation spectrum matches one of the lines
Xe: More stable, longer lifetime, flat emission in visible is sometimes beneficial

## Slide 22: Metal Halide Arc Lamp
Exfo, Intensilight, etc.

## Slide 23: Metal Halide Arc Lamp

## Slide 24
Benefit of Liquid Light Guide
Improved Illumination Uniformity

## Slide 25
Plasma Lamp
LIFI; Sutter XL
Electrodeless ‘arc’ source
Very bright
10,000 hr life

## Slide 26
Plasma Lamp
LIFI; Sutter XL

## Slide 27: LEDs
Good in the blue / red
Not so good in the green / yellow
Long lifetime
Fast switching

## Slide 28: Relative lamp power
Summary: Metal Halide best; LEDs still not quite bright enough.

## Slide 29
Matching Filters and Fluorophores
http://probes.invitrogen.com/resources/spectraviewer/
http://fluorescence.nexus-solutions.net/frames6.htm
https://www.omegafilters.com/curvo2/index.php

## Slide 30
Fluorescent molecules
Systems of conjugated bonds
that share electrons
Larger system  longer wavelength
C
C
C
C

## Slide 31
Excitation & emission maxima
Extinction coefficient 
 absorption cross section
 ≈ 50,000–100,000 M-1cm-1
Fluorescence quantum yield Qf
= # Photons emitted / # photons absorbed
Qf ≈ 25–90%
Brightness   Qf
Photo-bleaching quantum yield Qb
= average # of photons emitted
per molecule before bleaching.
Depends on environment.
 Qf / Qb
Parameters of fluorescent molecules
Iexc
l
C
: extinction
coefficent
(M-1cm-1)
l: length (cm)
C: concentration (M)
Iin
Iout

## Slide 32: Parameters for some common fluorophores
Dye lex lem e QY brightness
DAPI 350 470 27000 0.58 15.7
Fluorescein 490 520 67000 0.71 47.6
Alexa 488 494 517 73000 0.6 43.8
Rhodamine 554 573 85000 0.28 23.8
Cy3 554 568 130000 0.14 18.2
Cy5 652 672 200000 0.18 36
GFP 488 507 56000 0.6 33.6
mCherry 587 610 72000 0.22 15.8
CFP 433 475 32500 0.4 13
YFP 516 529 77000 0.76 58.5

## Slide 33
The Enemy:
Photo-bleaching
Decrease in emission intensity after exposure
Exciting a molecule once has a probability Qb of killing it
Each molecule will emit only a finite number of photons

## Slide 34
Photo-bleaching
Photostability varies between dyes

## Slide 35
Photo-bleaching of fluorescent proteins
mCherry
Single-exponential bleaching
Emerald
Double-exponential bleaching
Fast- and slow-bleaching populations?
Photons/molecue/s
Time
Photons/molecue/s
Time

## Slide 36
What to do about photo-bleaching?
Select fade-resistant dyes
Label densely
Decrease bleaching by anti-fade mounting media
  Glycerol
  Oxygen scavengers
  Free-radical scavengers
  Triplet state quenchers
Note: some anti-fade agents quench some dyes.
Budget the photons you have
  Only expose when observing
  Minimize exposure time & excitation power
  Use efficient filter combinations
  Use highly QE, low noise camera
  Use simple light path

## Slide 37: Effect of mounting medium on FITC bleaching
Ono et al. 2001, J. Histochem Cytochem. 49: 305-311

## Slide 38: Effect of mounting media on Alexa bleaching

## Slide 39: Fluorophore saturation
0
1
2
S1
S2
S0
hA
hA
Internal
conversion
hE
Fluorescence lifetime is ~1-5ns
Once illumination intensity is high enough to excite the fluorophore as soon as it deexcites, further intensity increases will not increase brightness
Usually only a problem for confocal

## Slide 40: Factors affecting overall brightness
Intrinisic brightness
Spectrum of arc lamp/lasers
Lamp/laser power
Filter set transmission
Quantum efficiency of detector
Photobleaching
Quenching / maturation / other dye-specific effects

## Slide 41: pH dependence of dyes
Mixed Fluorescein and Rhodamine
Rhodamine
Fluorescein
YFP variants

## Slide 42: Brief discussions of various fluors

## Slide 43: Traditional small molecule dyes

## Slide 44: Fluorescent dyes in Biology
FITC
Texas Red
Protein labeling: couple to amino- or sulfhydryl groups
Direct and indirect (immuno-) fluorescence
Notes: Show structure of FITC, TRITC, Cy3, Cy5
Discuss labeling strategies for proteins (targetting primary amino groups or sulfhydryls (cystein)
Explain direct and indirect immuno-fluorescence.
Show a number of images

## Slide 45
Fluorescent labeling
Direct labeling (& microinjection)
of target molecules
Direct immunofluorescence:
labeled antibodies against target
Indirect immunofluorescence:
Unlabeled antibodies against target
Labeled antibodies against those antibodies

## Slide 46: DNA Probes
Ethidium Bromide
~30 fold enhancement
DAPI
Hoechst 33258
Hoechst 33342
~20 fold enhancement
Notes: Show structure of Ethidium Bromide (intercalation), DAPI and Hoechst 33342 (minor groove binders).
Fluorescence emission is highly enhanced after DNA binding

## Slide 47: Other probes
ER-Tracker™ Blue-White DPX
MitoTracker Red CMXRos
Probes for Golgi, lysosomes, and peroxisomes are also available

## Slide 48: Small molecules – pros / cons
1000s available – huge spectral range
Easy to acquire
Precisely tailored properties, including environmental sensitivity
Require fixing and staining, which can lead to artifacts
Potential self-quenching and environmental sensitivity

## Slide 49
Example:
Newly synthesised connexins (ReAsH:Red) are added to the outer edges of existing gap junctions (FlAsH:Green). Gaietta et al 2002
FlAsH/ReAsH

## Slide 50: Quantum dots
“Artificial atoms” composed of small semiconductior nanocrystals

## Slide 51: Quantum dots - spectra

## Slide 52: Quantum dots – pros / cons
Little to no photobleaching
Very bright
Can use single excitation wavelength for multiple dyes
Narrow emission spectra
Large compared to small molecule dyes
Single quantum dots blink
Problems with non-specific binding

## Slide 53: Fluorescent Proteins and Genetically Encoded Tags See future lecture!
A. sulcata

## Slide 54: Resources
www.microscopyu.com
micro.magnet.fsu.edu
www.chroma.com (esp. their handbook on filter design)
www.probes.com (esp. their handbook/catalog)
Douglas B. Murphy “Fundamentals of Light Microscopy and Electronic Imaging”
James Pawley, Ed. “Handbook of Biological Confocal Microscopy, 3rd ed.”
Acknowledgements
Nico Stuurman / Mats Gustafsson / Mike Davidson

## Slide 55: Multi-photon excitation
0
1
2
S1
S0
hA
Internal
conversion
hE
hA
hA
Brad Amos, MRC, Cambridge