## Slide 1: Cameras For Microscopy
Ryan McGorty
26 March 2013

## Slide 2: Goals of this lecture
Understand what prevents the ideal camera from being a realty
Understand sources of noise in digital imaging
Decipher meaning of camera specification sheets
Understand basic architecture of digital sensors

## Slide 3: How image sensors sense an image
Photoelectric Effect
Hertz: observed in 1887
Einstein
For silicon, need ≈ 1.14 eV to release an electron
  So need λ < 1100 nm
1. Need way to detect photons

## Slide 4: Point Detectors
Photodiodes
  Photons will generate electron (photoelectron)
  But detecting few electrons is hard
  Need way to amplify the signal
Photon
e-
Notes: Photon generates electron, but few electrons are hard to detect

## Slide 5: Point Detectors
Avalanche Photodiode (APD)
  Initial e- becomes amplified through impact ionization
  Gains of 100 to 1000
Image from http://www.leica-microsystems.com/science-lab/sensors-for-true-confocal-scanning/

## Slide 6: Point Detectors
Photomultiplier Tube (PMT)
  Used widely in confocals
  High linearity
  Fast response
  High gain
Image from http://www.leica-microsystems.com/science-lab/sensors-for-true-confocal-scanning/

## Slide 7: How well are photons detected?
Quantum Efficiency
Notes: Higher gain doesn’t necessarily mean higher sensitivity

## Slide 8: How image sensors sense an image
2. Must detect photons over an array of pixels
CCD
CMOS

## Slide 9: How image sensors sense an image
3. Electrons generated in each pixel must be read out
CCD
CMOS
Serial devices where each pixel’s charge is read out one at a time
Each pixel contains amplifier electronics so read out can be much faster

## Slide 10: What is meant by “read out”?
Bit Depth | Number of gray levels
8 | 28 = 256
10 | 210 = 1024
12 | 212 = 4096
14 | 214 = 16384
16 | 216 = 65536
Photoelectrons
Voltage
10011101
Amplifier
Analog-Digital Converter (ADC)
× QE
Photons

## Slide 11: CCD Detectors
Image from http://www.microscopyu.com/articles/digitalimaging/ccdintro.html
Charge-Coupled Device
Invented at Bell Labs in 1970
2009 Nobel Prize in Physics
Widely used in TV, medical, astronomy cameras
Array of light sensitive MOS capacitors (pixels)
Incoming light generates electrons which are captured in a potential well
Electrodes, or gates, move the charge

## Slide 12: Front or Back Illuminated
From http://learn.hamamatsu.com/articles/quantumefficiency.html
Traditional, front-illuminated have wiring in front of photosensitive region
  This blocks some light, reducing QE
Back-illuminated CCDs:
  Back side of the CCD is etched to 10-15 microns
  More fragile and costly, but higher QE
Retina
Nerves
Optic nerve
Vertebrate
Octopus
http://en.wikipedia.org/wiki/Cephalopod_eye

## Slide 13: CCD Detectors
e-

## Slide 14: CCD Detectors
Photoelectrons from each pixel must be shuttled to the readout area to be amplified, converted to a voltage and then a digital number.
Three-phase clocking method

## Slide 15: Bucket brigade analogy for read out
Rain accumulates in buckets
Rows of buckets shifted to readout row
Readout row shifted bucket-by-bucket to measuring device
Unless rain stops, last bucket to be read out will have much more water than the first

## Slide 16: Transferring photoelectrons out
Full frame
  Full area of CCD collects light
  During readout, shutter must block light
Incoming light
Shutter

## Slide 17: Transferring photoelectrons out
Full frame
Frame transfer
  Half of CCD collects light; other half is for storage
  After exposure, charge transferred to storage half and read out
Incoming light
Light blocking mask

## Slide 18: Transferring photoelectrons out
Full frame
Frame transfer
Interline transfer
  Alternating columns of light sensitive and storage/transfer pixels
  Photosensitive pixels collect charge; charge shifted to neighboring columns and read out
Incoming light

## Slide 19: Transferring photoelectrons out
Pros:
Nearly all of CCD area is usable
Cons:
Frame rates limited by shutter speed
Pros:
Fast: exposure and readout occur simultaneously
Cons:
Half of CCD is not light sensitive
Pros:
Fast: shift of only one pixel needed to store charge
Cons:
Reduced fill factor (abated with microlenses)
Full frame
Frame-transfer
Interline
Readout register
Output amplifier

## Slide 20: CMOS Detectors
Image from http://www.olympusmicro.com/primer/digitalimaging/cmosimagesensors.html
Complementary Metal Oxide Semiconductor
Transistors in each pixel convert charge to voltage
More can be done within a pixel meaning frame read out can be faster
Fabricated much like microprocessors and RAM so are cheaper to make
Used in webcams, phone cameras since they use less power

## Slide 21: CMOS read out
Each pixel has at least 3 transistors
  For converting charge to voltage; resetting pixel; transferring voltage
Voltage transferred along columns to amplifiers and ADCs
Each row is then read out to memory
ADC
ADC
ADC
ADC

## Slide 22: CMOS read out
Rolling Shutter:
  Exposure start is delayed between each row
  Readout goes row by row
  Offset time between rows as low as 10 μs
Global vs. Rolling Shutter
Exposure starts
Exposure starts
Exposure starts
ADC
ADC
ADC
ADC

## Slide 23: CMOS read out
Rolling Shutter:
  Exposure start is delayed between each row
  Readout goes row by row
  Offset time between rows as low as 10 μs
Global vs. Rolling Shutter
Readout starts
Readout starts
Readout starts
ADC
ADC
ADC
ADC

## Slide 24: CMOS read out
Rolling Shutter:
  Exposure start is delayed between each row
  Readout goes row by row
  Offset time between rows as low as 10 μs
Global vs. Rolling Shutter
http://www.andor.com/pdfs/literature/Tech_Note_Rolling_Global_Shutter.pdf

## Slide 25: CMOS read out
Rolling Shutter:
  Exposure start is delayed between each row
  Readout goes row by row
  Offset time between rows as low as 10 μs
Global Shutter:
  All pixels begin and end exposure at same time
  Requires extra transistor in each pixel
  Higher read noise
  Half the maximum frame rate
Global vs. Rolling Shutter

## Slide 26: Many options
CCD
CMOS

## Slide 27: Camera Spec Sheets

## Slide 28: The Ideal Camera
High resolution
  Optically resolvable features span a couple pixels
Large area
  Covers field of view
Linear response
  Twice the photons gives twice the output
Uniform response
  Same response no matter where photon lands
Fast
  Can capture dynamic processes
High sensitivity
  All photons get measured

## Slide 29: Going into the spec sheet
SONY Interline CCD chip
Notes: Camera companies don’t generally make the chip sensor, only the electronics that surround it

## Slide 30: Going into the spec sheet
Suppose you want 3 pixels per resolvable unit of 300 nm:
0.3 μm × M = 3 × 6.45 μm
M = 64.5×

## Slide 31: Going into the spec sheet

## Slide 32: How well are photons detected?
Quantum Efficiency

## Slide 33: Going into the spec sheet

## Slide 34: NOISE!
What is noise?
Uncertainty in measured signal
Looks like random fluctuations in intensity

## Slide 35: NOISE!
The different types of noise
  Fixed Pattern Noise
    Pixels respond to light or photoelectrons differently
      [Perhaps because piece of dirt in front of some]
    [PN is the FPN quality factor]
    Correctable through process of flat fielding
    Can be an issue with CMOS since electronics associated with each pixel can be different

## Slide 36: NOISE!
The different types of noise
  Fixed Pattern Noise
  Dark Noise
    Temperature dependent; reduced by cooling camera
  Clock-induced Charge Noise
    As charge gets moved around, new electrons can be generated

## Slide 37: NOISE!
The different types of noise
  Fixed Pattern Noise
  Dark Noise
  Clock-induced Charge Noise
  Shot Noise
    Photons are discrete; follow Poisson distribution
    Photon Shot Noise
Average of 1 photon per pixel
Photon distribution on 1 million pixels
Average of 5 photons per pixel
10 photons/pixel

## Slide 38: NOISE!
The different types of noise
  Fixed Pattern Noise
  Dark Noise
  Clock-induced Charge Noise
  Shot Noise
  Read Noise
    Not a function of signal. Due to electronics converting charge to voltage
    Generally increases with increased readout speed
    Sets the minimum signal that can be detected
Decreasing read noise
Same signal.

## Slide 39: NOISE!
The different types of noise
  Fixed Pattern Noise
  Dark Noise
  Clock-induced Charge Noise
  Shot Noise
  Read Noise
Signal to Noise Ratio

## Slide 40: NOISE!
Signal to Noise Ratio
Image from Photon Transfer by James R. Janesick

## Slide 41: How do you optimize SNR?
Choose a camera that has:
  High QE
    Back illuminated
  Low read noise
  Cooling ability
  Spatial uniformity
When imaging:
  Go at slowest speeds
  Use largest pixel size
    Binning
Binning increases the SNR by a factor of the sqrt of number of pixels binned

## Slide 42: Nearly eliminating read noise: EMCCDs
Amplifies signal to push it well above read noise
In CCDs: rows gets transferred to edge to readout register
EMCCDs add an electron multiplication (EM) register
As charge is transferred through EM register, impact ionization increases charge on order of 100-1000×
Image from http://www.andor.com/learning-academy/electron-multiplying-ccd-cameras-the-technology-behind-emccds

## Slide 43: Noise in EMCCDs
Read noise = 10 e-
Read noise = 0.5 e-
Read noise = 0.5 e- Plus EM associated noise, Fn
Notes: Simulation details: On av 300 photons per molecule. 10 photons/pixel background

## Slide 44: A new option for when photons are few: sCMOS
“scientific” CMOS
  Latest generation of CMOS cameras
  Sensor split into top and bottom halves
    Both are read out independently to increase speed
  Each column as dual amplifiers and ADC
    For high and low gain readout to maximize dynamic range

## Slide 45: NOISE!
Photoelectrons
Voltage
10011101
Amplifier
Analog-Digital Converter (ADC)
× QE
Photons
Shot noise
Readout noise
Fixed pattern noise
EM Gain
Clock induced charge noise
EM noise factor
Dark noise

## Slide 46: Measuring noise
Steps to generate photon transfer curve:
  Record several sequences of frames at various illumination levels starting with completely dark
  Use homogenous illumination,
  Or correct for inhomogeneity as well as fixed pattern noise by subtracted neighboring frames
  Plot standard deviation of signal versus mean signal

## Slide 47
Noise includes shot noise and read noise. Fixed pattern noise eliminated by taking the difference of consecutive frames.
Slope of ½
Slope of 0
sCMOS

## Slide 48
Images show variance on a log scale
Slope of 0
sCMOS

## Slide 49
Slope of ½
Slope of 1
Fixed pattern noise is now included.
sCMOS

## Slide 50
Images show average of 400 frames
sCMOS

## Slide 51
Read Noise
Shot Noise
Fixed Pattern Noise
sCMOS
By looking at shot noise vs. signal: ADU conversion is about 2.2 DN/e-
Read noise: noise vs. signal levels out around 3.8 DN = 1.7e-
Fixed pattern noise: about 1.7% of signal

## Slide 52
sCMOS
Read noise: noise vs. signal levels out around 1.7e-

## Slide 53
By looking at shot noise vs. signal: ADU conversion is about 1.0 DN/e-
Read noise: noise vs. signal levels out around 5 DN = 5 e-
Shot Noise
Read Noise
Full well regime
CCD
Fixed pattern noise and illumination inhomogeneity eliminated by taking the difference of consecutive frames.

## Slide 54
EMCCD
ADU ≈ 0.1 DN/e-
ADU ≈ 5 DN/e-
ADU ≈ 10 DN/e-

## Slide 55
EMCCD
Read noise ≈ 30 e-
Read noise ≈ 0.9 e-
Read noise ≈ 0.5 e-
When plotted in units of e- rather than DN benefit of EM gain can be seen: read noise is effectively reduced by the amount of gain

## Slide 56
All three cameras above read noise
Cross over between EMCCD and sCMOS

## Slide 57: Going back to the spec sheet
Nr = 6 e-
QE ≈ 45% at ≈ 700nm

## Slide 58: Going back to the spec sheet
Dark current: 1 e-/pixel/s @ 23 °C

## Slide 59: Going back to the spec sheet
Dynamic range: 2,667:1
Read noise = 6 e-
Full well capacity = 16,000 e-
Dynamic range = 16,000/6 = 2667
Dynamic range A/D = 14 bit
Gray levels = 214 = 16,384

## Slide 60: Thanks!
For more information:
  http://learn.hamamatsu.com/articles
  http://www.microscopyu.com/articles/digitalimaging/index.html/
  Photon Transfer by James Janesick