## Slide 1: Optical Sectioning and Confocal Microscopy Kurt Thorn

## Slide 2: The goal: build 3D images of biological samples
Embryonic mouse lung; Nan Tang, Martin lab

## Slide 3: Optical Sectioning and 3D reconstruction
x
y
z

## Slide 4: Optical Sectioning and 3D reconstruction
x
y
z

## Slide 5: Optical Sectioning and 3D reconstruction
x
y
z
=

## Slide 6: A series of optical sections
Yeast mitochondria; Susanne Rafelski

## Slide 7
Fluorescence Illumination of a single point
Sample
Objective lens
Excitation light
Tube lens
Emission light
Problem – fluorescence is emitted along entire illuminated cone, not just at focus
Camera

## Slide 8: Widefield fluorescence imaging
Widefield Illumination
Point Illumination

## Slide 9
Measured
x
z
Calculated
The Point Spread Function (PSF)
Z=0
Z=+2µm
Z=-2µm
2D PSF
for different defocus
The image of a point object
x
y
3D PSF

## Slide 10: The Problem of Out-of-Focus Light
Prevent it from reaching the detector
  Confocal microscopy
Prevent it from being excited
  Two-photon, light sheet microscopy, TIRF
Remove it after the fact (computationally)
  Deconvolution
Three Options

## Slide 11
Fluorescence Illumination of a single point
Sample
Objective lens
Excitation light
Tube lens
Emission light
Problem – fluorescence is emitted along entire illuminated cone, not just at focus
Camera

## Slide 12
The confocal microscope
Sample
Objective lens
Excitation light
Tube lens
Emission light
Pinhole
Detector
Scan excitation spot point-by-point to build up image

## Slide 13: How big should your pinhole be?
Resolution is limited by the point-spread function
Airy disk radius
≈ 0.61  /NA
Y
Z
X

## Slide 14: How big should your pinhole be?
Want pinhole to pass entire Airy disk
Airy disk diameter
≈ 1.22  /NA
Y
X
Width of point spread function at pinhole: Airy disk diameter × magnification of lens

## Slide 15: How big should your pinhole be?
Width of point spread function at pinhole = Airy disk diameter × magnification of lens = 1 Airy unit = resolution of lens × magnification of lens × 2
  100x / 1.4 NA: resolution = 220nm, so 1 Airy unit = 44 mm
  40x / 1.3 NA: resolution = 235nm, so 1 Airy unit = 19 mm
  20x / 0.75 NA: resolution = 407nm, so 1 Airy unit = 16 mm
  10x / 0.45 NA: resolution = 678nm, so 1 Airy unit = 14 mm

## Slide 16: Light sources
Sample
Objective lens
Excitation light
Excitation light must be focused to a diffraction limited spot
Could be done with an arc lamp and pinhole – but very inefficient
Enter the laser:
Perfectly collimated and high power

## Slide 17
Scanning
Sample
Objective lens
Changing entrance angle of illumination moves illumination spot on sample
The emission spot moves, so we have to make sure pinhole is coincident with it

## Slide 18
Confocal optical path
Objective
x
y
Relay lens
Laser
Scanning mirrors
Dichroic
Pinhole
Detector

## Slide 19: What do you get?

## Slide 20: Detectors - PMTs
Must be fast – confocal beam spends only a few ms on each pixel
  Photomultiplier tubes
Pulse width for single photon ~ 10-100ns
Very linear
Very high gain
~ 0 read noise

## Slide 21: Detectors - PMTs
Peak QE for standard PMT: ~30%
Front-illuminated CCD: ~60%
Back-illuminated CCD: ~95%
GaAsP PMTs: ~40-50%
However, GaAsP PMTs can be damaged by exposure to too much light.

## Slide 22: Detectors - APDs
Photons produce photoelectrons
Photoelectrons are amplified by collisional ejection of electrons
Small dynamic range

## Slide 23: Hybrid Detectors
GaAsP photocathode bombards an APD
Very high gain
Excellent for single photon counting
Damaged by excessive light

## Slide 24: Confocal vs. Widefield
Confocal
Widefield
Tissue culture cell with 60x / 1.4NA objective

## Slide 25: Confocal vs. Widefield
Confocal
Widefield
20 mm rat intestine section recorded with 60x / 1.4NA objective

## Slide 26: Doing more with your photons
Spectral Detection
Prism + Adjustable Slits: Leica SP microscopes

## Slide 27: Doing more with your photons
Spectral Detection
Prism + PMT-array: Zeiss and Nikon

## Slide 28: Doing more with your photons
Time-Correlated Single Photon Counting
Single-photon detector
Pulsed Laser

## Slide 29
The confocal microscope: Drawbacks
Sample
Objective lens
Excitation light
Tube lens
Emission light
Pinhole
Detector
Scan excitation spot point-by-point to build up image
Problems:
Slow (~1 sec to acquire an image)
Low light efficiency (due to use of PMT as detector)
Solution:
Use multiple pinholes and a camera

## Slide 30: A Solution: Spinning Disk Confocal
Image with many pinholes at once, so fast
Use CCD as detector, so much higher QE

## Slide 31: The downside to the spinning disk

## Slide 32: The downside to the spinning disk

## Slide 33: The downside to the spinning disk

## Slide 34: The downside to the spinning disk
Limited out of focus rejection

## Slide 35: Pros/Cons of spinning disk
Fast – multiple points are illuminated at once
Photon efficient – high QE of CCD
Gentler on live samples – usually lower laser power
Fixed pinhole size (usually)
Small field of view (usually)
Crosstalk through adjacent pinholes limits sample thickness

## Slide 36: Out-of-focus rejection
Defocus
T
Wolleschensky et al. J. Biomed. Opt. 11(6) 2006

## Slide 37: Examples
Drosophila S2 cell expressing GFP-H2B and mCherry-tubulin (Nico Stuurman and Ron Vale)
S. cerevisiae expressing a mitochondrially targeted RFP, Susanne Rafelski, Marshall lab

## Slide 38: When to use confocal?
Confocal is not a magic bullet
  It is extremely wasteful of photons
  Laser-scanning confocal is 100 – 200-fold less sensitive than widefield
  Spinning-disk confocal is ~4-fold less sensitive than widefield
Murray JM et al., J. Microsc. 2007, 228: p.390-405

## Slide 39: When to use confocal?
Confocal is not a magic bullet
  It is extremely wasteful of photons
  High laser power generally result in more photobleaching and photodamage.
For thin specimens, widefield epifluorescence is better – especially with deconvolution
Confocal excels with thick, heavily stained specimens

## Slide 40: When to use confocal?
What is thick?
  A good rule of thumb is 10 × the depth of field of the objective
  100x / 1.4 NA: d.o.f. 0.66mm
  20x / 0.75 NA: d.o.f 2.3 mm
Sample preparation is KEY for imaging thick specimens
  Confocal does not fix scattering, refractive index mismatch, or everything else that can go wrong – it only removes out of focus light

## Slide 41: Sample preparation
For fixed samples: match refractive index of mounting media to immersion oil.
  Mount in immersion oil itself, BABB, benzyl alcohol/glycerol, 2,2’-thiodiethanol, or other high-RI mounting medium
Clearing to remove lipids and other scattering substances is also important
For live samples, use water immersion lenses

## Slide 42: Sample Preparation
Samples imaged with 20x / 0.75 air objective on spectral confocal
Sections acquired ~ 50 mm into tissue
Embryonic mouse lungs; samples from Nan Tang, Martin Lab

## Slide 43: Sample Preparation
Samples imaged with 40x / 1.3 oil objective on spectral confocal
Sections acquired ~ 50 mm into tissue
Embryonic mouse lungs; samples from Nan Tang, Martin Lab

## Slide 44: Clearing and mounting summary
Both clearing and refractive index matching are important.
BABB and methyl salicylate clear very well and give best image depth, but may disrupt cell morphology
TDE preserves cell morphology and also allows for decent imaging depth

## Slide 45: Newer Clearing Techniques
Scale
ScaleA2: 4M urea, 10% glycerol, 0.1% Triton X-100
Hama et al. Nat. Neurosci. 14, p. 1481 (2011)

## Slide 46: Newer Clearing Techniques
CLARITY
Chung et al. Nature 2013; http://clarityresourcecenter.org/

## Slide 47: Resources
http://www.microscopyu.com
http://micro.magnet.fsu.edu
James Pawley, Ed. “Handbook of Biological Confocal Microscopy, 3rd ed.”
http://www.hamamatsu.com/resources/pdf/etd/PMT_handbook_v3aE.pdf
Acknowledgements
Steve Ross, Mats Gustafsson