## Slide 1: Principles and Practices of Light Microscopy II: Brightfield optics, resolution, and aberrations

## Slide 2: Trans-illumination Microscope
Sample
Objective
Tube lens
Intermediate image plane
Projection Eyepiece
Object plane
Back focal plane (pupil)
Secondary pupil plane
Camera
Final image plane
Imaging
path
Aperture iris
Field iris
(pupil plane)
(image plane)
(pupil plane)
Light source
Illumination
path
Collector
Condenser lens
Field lens
The aperture iris
controls the range of
illumination angles
The field iris
controls the
illuminated
field of view

## Slide 3: By far the most important part: the Objective Lens
Each major manufacturer sells 20-30 different categories of objectives.
What are the important distinctions?

## Slide 4: Working Distance
In general, high NA lenses have short working distances
However, extra-long working distance objectives do exist
Some examples:
10x/0.3 WD = 15.2mm
20x/0.75 WD = 1.0mm
100x/1.4 WD = 0.13mm

## Slide 5: The focal length of a lens depends on the refractive index…
Focal length f
f  1/(n-1)
Refractive index n

## Slide 6: … and the refractive index depends on the wavelength (“dispersion”)
Glass
types

## Slide 7:  Chromatic aberration
Different colors get focused to different planes
Not good…

## Slide 8: Dispersion vs. refractive index of different glass types
Abbe dispersion number
Refractive
index
(Higher dispersion)

## Slide 9: Achromatic Lenses
Use a weak negative flint glass element
to compensate the dispersion of a
positive crown glass element

## Slide 10: Achromats and Apochromats
Achromat (2 glass types)
Apochromat (3 glass types)
Wavelength
Focal length
Simple lens

## Slide 11: Correction classes of objectives
Achromat (cheap)
Fluor “semi-apo” (good correction, high UV transmission)
Apochromat (best correction)
Correction for other (i.e. monochromatic) aberrations also improves in the same order

## Slide 12
Using the wrong immersion oil can induce axial chromatic aberration
(X-Y)
(X-Z)
Cargille
Steve Ross
Notes: All taken with TYPE A oil. 60x/1.40. Different dispersions. Lateral is an X-Y projection. Axial is 3D in Z. NF and DF oils have some auto fluorescence in visible but not near UV and UV. Use type A for visible fluorescence.

## Slide 13: Putting one brand of objectives onto another brand of microscope?
Usually a bad idea:
May not even fit
May get different magnification than
is printed on the objective
Incompatible ways of
correcting lateral chromatic
aberration (LCA)
 mixing brands can produce severe LCA
Pitch = 0.75
Tube lens focal length
Nikon 200
Leica 200
Olympus 180
Zeiss 165
LCA correction:
In objective In tube lens
Nikon Leica Olympus Zeiss

## Slide 14: Lateral chromatic aberration (= LCA, lateral color, chromatic difference of magnification)
= Different magnification for different colors
Object
Image

## Slide 15: Curvature of Field
Focal plane
Focal plane Focal surface
sample
Focal
surface
objective
Tube lens

## Slide 16: Plan objectives
Corrected for field curvature
More complex design
Needed for most photomicrography
Plan-Apochromats have the highest performance
(and highest complexity and price)

## Slide 17
Aberrations
Chromatic aberrations
  Longitudinal chr. Ab.
  Lateral chr. Ab.
Curvature of field
Distortion
Wavefront aberrations
  Spherical aberration
  Astigmatism
  Coma
  …

## Slide 18
Geometric Distortion
= Radially varying magnification
Barrel
distortion
Pincushion
distortion
Object
Image
May be introduced by the projection eyepiece

## Slide 19: Diffraction by an aperture
Larger aperture

weaker diffraction
Light spreads to new angles
drawn as waves

## Slide 20: Diffraction by an aperture
The pure, “far-field”
diffraction pattern
is formed at  distance…
…or can be formed
at a finite distance
by a lens…
…as happens in a microscope
Objective pupil
Intermediate
image
Tube lens
drawn as rays

## Slide 21: The Airy Pattern
= the far-field diffraction pattern from a round aperture
“Airy disk” diameter
d = 2.44  f/d
(for small angles d/f)
d
f
Height of
first ring
 1.75%

## Slide 22: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Diffraction spot
on image plane
= Point Spread Function
Sample
Objective

## Slide 23: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Sample
Objective
Diffraction spot
on image plane
= Point Spread Function

## Slide 24: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Sample
Objective
Diffraction spot
on image plane
= Point Spread Function

## Slide 25: Aperture and Resolution
Image resolution improves with aperture size
Sample
Objective
Tube lens
Back focal plane aperture
Intermediate image plane
Diffraction spot
on image plane
(resolution)

Numerical Aperture (NA)
NA = n sin()
= light gathering angle
n = refractive index of sample
where:

## Slide 26: Numerical Aperture
4X / 0.20 NA
 = 11.5°
100X / 0.95 NA
 = 71.8°

## Slide 27: Numerical Aperture
Numerical Aperture:
NA = n sin()
Compare:
Snell’s law:
n1 sin(1) = n2 sin(2)
1
2

Cover
glass
Sample
NA cannot exceed
the lowest n between the
sample and the objective lens
n sin() doesn’t change at horizontal interfaces
sin(anything)  1

## Slide 28: Numerical Aperture
Numerical Aperture:
NA = n sin()
Compare:
Snell’s law:
n1 sin(1) = n2 sin(2)
1
2
n sin() doesn’t change at horizontal interfaces
sin(anything)  1
NA cannot exceed
the lowest n between the
sample and the objective lens
NA >1 requires fluid immersion

Cover
glass
Sample
Immersion fluid

## Slide 29: Immersion Objectives
Oil immersion:
  n  1.515
  max NA  1.4 (1.45–1.49 for TIRF)
Glycerol immersion:
  n  1.45 (85%)
  max NA  1.35 (Leica)
Water immersion:
  n  1.33
  max NA  1.2
NA can approach
the index of the
immersion fluid
NA cannot exceed
the lowest n between the
sample and the objective lens
NA >1 requires fluid immersion

## Slide 30
Resolution
Ernst Abbe’s argument (1873)
Sample
Objective lens
Back focal
plane
Condenser
Light source
Consider a striped sample ≈ a diffraction grating
b
Diffracted beams
d sin(b) = l
d
Consider first
a point light source
Smaller d
 larger b
If b > a, only one spot makes it through
 no interference  no image formed
Resolution (smallest resolvable d):
dmin = lsample/sin(a) = l/n sin(a) = l/NA

## Slide 31
(Abbe’s argument, continued)
Now consider oblique illumination
(an off-axis source point):
bout
d [sin(bin) + sin(bout) ] = l
One spot hopelessly lost,
but two spots get through
 interference  image formed!
Resolution (smallest resolvable d)
with incoherent illumination (all possible illumination directions):
dmin = l/(NAobj + NAcondenser )
l/2 NA if NAcondenser  NAobj (“Filling the back focal plane”)
bin
Two spots get through if
bout < a and bin < a.
d

## Slide 32
aobjective
For the highest resolution,
we need to have
condenser  objective
NAcondenser  NAobjective
with oil immersion objectives,
we need an oil immersion condenser!
acondenser
Filling the back focal plane
In trans-illumination microscopy, to get maximum resolution,
the illumination must “fill the back focal plane”
Objective
Condenser
Light source
Back
focal
plane

## Slide 33
The Condenser
Tasks:
  Illuminate at all angles < aobjective
  Concentrate light on the field of view for all objectives to be used
Problem:
  Low mag objectives have large FOV,
  High mag objectives have large a
  (With 2X and 100x objectives we need (100/2)2 = 2500 times more light than any objective uses!)
Solutions:
  Compromise
  Exchangable condensers,
  swing-out front lenses,…
Grade of correction
NA

## Slide 34
Aperture, Resolution & Contrast
Can adjust the condenser NA with the aperture iris
Sample
Objective
Tube lens
Imaging
path
Aperture iris
Field iris
Light source
Illumination
path
Collector
Condenser lens
Field lens
Back aperture
Intermed. image
Q: Don’t we always want
it full open??
A: No
Why? Tradeoff:
resolution vs. contrast

## Slide 35: Numerical Aperture and Resolution
Snell’s law:
n1 sin(1) = n2 sin(2)
1
2
Cover
glass
Sample
Immersion
fluid

## Slide 36: Numerical Aperture and Resolution
Cover
glass
Sample
Immersion
fluid
Sample, coverglass, immersion fluid, and top lens of objective all have same refractive index.
What happens if we change that refractive index?
Resolution improves with RI! Why?

## Slide 37: Light travels more slowly in matter
n = 1
n > 1
n = 1
And the wavelength shortens
Recall Abbe’s experiment: dmin = lsample/sin(a) = l/n sin(a) = l/NA

## Slide 38: Modifying RI to change resolution
Diamond solid immersion lens
Appl. Phys. Lett. 97, 241902 (2010)
Olympus 1.65 NA TIRF lens: Uses sapphire coverslips (n=1.76), diiodomethane (n=1.74)

## Slide 39: Why not use NA 1.65 lenses?
Water
n = 1.33
Glass n = 1.51
What’s the largest angle we can collect?
NA 1.4 oil immersion objective
Water
n = 1.33
Sapphire n = 1.78
NA 1.65 oil immersion objective

## Slide 40: Effective NA is limited by the sample
To use a high NA objective, need a high RI mounting medium
Water 1.33 Glycerol 1.45 Vectashield 1.44 Prolong Gold 1.39 – 1.46 2,2-thiodiethanol 1.52 Methyl Salicylate 1.53 Benzyl benzoate 1.57

## Slide 41: NA and Resolution
High NA Objective Low NA Objective

## Slide 42
Alternate Definitions of Resolution
As the Full Width at Half Max
(FWHM) of the PSF
As the diameter of the Airy disk
(first dark ring of the PSF)
= “Rayleigh criterion”
(Probably most common definition)
Airy disk radius
≈ 0.61  /NA
FWHM
≈ 0.353  /NA

## Slide 43: Objective Types
Field flatness
Plan or not
Phase rings for phase contrast
Positive or negative
Diameter of ring (number)
Special Properties
Strain free for Polarization or DIC
Features
Correction collar for spherical aberration
Iris
Spring-loaded front end
Lockable front end
Basic properties
Magnification
Numerical Aperture (NA)
Infinite or finite conjugate
Cover slip thickness if any
Immersion fluid if any
Correction class
Achromat
Fluor
Apochromat

## Slide 44: Further reading
www.microscopyu.com
micro.magnet.fsu.edu
Michael W. Davidson and Douglas B. Murphy “Fundamentals of Light Microscopy and Electronic Imaging”
James Pawley, Ed. “Handbook of Biological Confocal Microscopy, 3rd ed.”
Acknowledgements
Ron Vale / Mats Gustafsson / Steve Ross