## Slide 1
(Image: T. Wittman, Scripps)
Introduction to
Light Microscopy

## Slide 2
The Light Microscope
Four centuries of history
Vibrant current development
One of the most widely used research tools

## Slide 3
A. Khodjakov et al.

## Slide 4: Major Imaging Functions of the Microscope
Magnify
Resolve features
Generate Contrast
Capture and Display Images

## Slide 5
An Upright Epifluorescence Microscope

## Slide 6: Light travels more slowly in matter
n = 1
n > 1
n = 1
v = c/n
The speed ratio is the Index of Refraction, n

## Slide 7: Refractive Index Examples
Depends on wavelength and temperature
Vacuum 1
Air 1.0003
Water 1.333
Cytoplasm 1.35–1.38 ?
Glycerol 1.475 (anhydrous)
Immersion oil 1.515
Fused silica 1.46
Optical glasses 1.5–1.9
Diamond 2.417

## Slide 8: Lenses work by refraction
Incident light
focus
Focal length f

## Slide 9: Finite vs. Infinite Conjugate Imaging
Object
f0
f0
f0
Object
f0
>f0
f0
Image
Finite conjugate imaging (older objectives)
Image at infinity
Infinite conjugate imaging (modern objectives).
f1
f1
Magnification:
Image
(uncritical)
 Need a tube lens

## Slide 10: Back focal plane
Object
f0
f0
f0
Back focal plane
Rays that leave the object with the same angle
meet in the objective’s back focal plane

## Slide 11: The Compound Microscope
Sample
Objective
Tube lens
Primary or intermediate
image plane
Eyepiece
Back focal plane (pupil)
Exit pupil
Object plane

## Slide 12: The Compound Microscope
Sample
Objective
Tube lens
Intermediate image plane
Eyepiece
Object plane
Back focal plane (pupil)
Exit pupil
Eye
Final image

## Slide 13: The Compound Microscope
Sample
Objective
Tube lens
Intermediate image plane
Eyepiece
Object plane
Back focal plane (pupil)
Exit pupil
Eye
Final image

## Slide 14: The Compound Microscope
Sample
Objective
Tube lens
Intermediate image plane
Eyepiece
Object plane
Back focal plane (pupil)
Exit pupil
Eye
Final image

## Slide 15: The Compound Microscope
Sample
Objective
Tube lens
Intermediate image plane
Eyepiece
Object plane
Back focal plane (pupil)
Exit pupil
Eye
Final image

## Slide 16: The Compound Microscope
Sample
Objective
Tube lens
Intermediate image plane
Projection Eyepiece
Object plane
Back focal plane (pupil)
Secondary pupil
Camera
Final image

## Slide 17: Eyepieces (Oculars)
Features
Magnification (10x typical)
“High eye point” (exit pupil high
enough to allow eyeglasses)
Diopter adjust (at least one
must have this)
Reticle or fitting for one
Eye cups

## Slide 18: Trans-illumination Microscope
Sample
Objective
Tube lens
Intermediate image plane
Projection Eyepiece
Object plane
Back focal plane (pupil)
Secondary pupil plane
Camera
Final image plane
Imaging
path
Aperture iris
Field iris
(pupil plane)
(image plane)
(pupil plane)
Light source
Illumination
path
Collector
Condenser lens
Field lens
The aperture iris
controls the range of
illumination angles
The field iris
controls the
illuminated
field of view

## Slide 19: Köhler Illumination
Object plane
(pupil plane)
(image plane)
(pupil plane)
Sample
Aperture iris
Field iris
Light source
Critical Illumination
Each light source point produces a parallel beam of light at the sample
Uniform light intensity at the sample even if the light source is “ugly” (e.g. a filament)
The source is imaged onto the sample
Usable only if the light source is perfectly uniform

## Slide 20: Conjugate Planes in A Research Microscope

## Slide 21
Two ways:
“Eyepiece telescope”
“Bertrand lens”
How view the pupil planes?

## Slide 22: By far the most important part: the Objective Lens
Each major manufacturer sells 20-30 different categories of objectives.
What are the important distinctions?

## Slide 23: Working Distance
In general, high NA lenses have short working distances
However, extra-long working distance objectives do exist
Some examples:
10x/0.3 WD = 15.2mm
20x/0.75 WD = 1.0mm
100x/1.4 WD = 0.13mm

## Slide 24: The focal length of a lens depends on the refractive index…
Focal length f
f  1/(n-1)
Refractive index n

## Slide 25: … and the refractive index depends on the wavelength (“dispersion”)
Glass
types

## Slide 26:  Chromatic aberration
Different colors get focused to different planes
Not good…

## Slide 27: Dispersion vs. refractive index of different glass types
Abbe dispersion number
Refractive
index
(Higher dispersion)

## Slide 28: Achromatic Lenses
Use a weak negative flint glass element
to compensate the dispersion of a
positive crown glass element

## Slide 29: Achromats and Apochromats
Achromat (2 glass types)
Apochromat (3 glass types)
Wavelength
Focal length error
Simple lens

## Slide 30: Correction classes of objectives
Achromat (cheap)
Fluor “semi-apo” (good correction, high UV transmission)
Apochromat (best correction)
Correction for other (i.e. monochromatic) aberrations also improves in the same order

## Slide 31: Curvature of Field
Focal plane
Focal plane Focal surface
sample
Focal
surface
objective
Tube lens

## Slide 32: Plan objectives
Corrected for field curvature
More complex design
Needed for most photomicrography
Plan-Apochromats have the highest performance
(and highest complexity and price)

## Slide 33: What is the Resolution Limit of an Objective?
i.e. what is the smallest separation between two objects you can detect?

## Slide 34: Diffraction by a periodic structure (grating)

## Slide 35: Diffraction by a periodic structure (grating)

In phase if:
d Sin() = m 
for some integer m
d
d Sin()


## Slide 36: Diffraction by an aperture
See “Teaching Waves with Google Earth” http://arxiv.org/pdf/1201.0001v1.pdf for more

## Slide 37: Diffraction by an aperture
Larger aperture

weaker diffraction
Light spreads to new angles
drawn as waves

## Slide 38: Diffraction by an aperture
The pure, “far-field”
diffraction pattern
is formed at  distance…
…or can be formed
at a finite distance
by a lens…
…as happens in a microscope
Objective pupil
Intermediate
image
Tube lens
drawn as rays

## Slide 39: The Airy Pattern
= the far-field diffraction pattern from a round aperture
“Airy disk” diameter
d = 2.44  f/d
(for small angles d/f)
d
f
Height of
first ring
 1.75%

## Slide 40: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Diffraction spot
on image plane
= Point Spread Function
Sample
Objective

## Slide 41: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Sample
Objective
Diffraction spot
on image plane
= Point Spread Function

## Slide 42: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Sample
Objective
Diffraction spot
on image plane
= Point Spread Function

## Slide 43: Aperture and Resolution
Image resolution improves with aperture size
Sample
Objective
Tube lens
Back focal plane aperture
Intermediate image plane
Diffraction spot
on image plane
(resolution)

Numerical Aperture (NA)
NA = n sin()
= light gathering angle
n = refractive index of sample
where:

## Slide 44: Numerical Aperture
4X / 0.20 NA
 = 11.5°
100X / 0.95 NA
 = 71.8°

## Slide 45: Immersion Objectives
Oil immersion:
  n  1.515
  max NA  1.4 (1.45–1.49 for TIRF)
Glycerol immersion:
  n  1.45 (85%)
  max NA  1.35 (Leica)
Water immersion:
  n  1.33
  max NA  1.2
NA can approach
the index of the
immersion fluid
NA cannot exceed
the lowest n between the
sample and the objective lens
NA >1 requires fluid immersion

## Slide 46
Resolution
Ernst Abbe’s argument (1873)
Sample
Objective lens
Back focal
plane
Condenser
Light source
Consider a striped sample ≈ a diffraction grating
b
Diffracted beams
d sin(b) = l
d
Consider first
a point light source
Smaller d
 larger b
If b > a, only one spot makes it through
 no interference  no image formed
Resolution (smallest resolvable d):
dmin = lsample/sin(a) = l/n sin(a) = l/NA

## Slide 47
(Abbe’s argument, continued)
Now consider oblique illumination
(an off-axis source point):
bout
d [sin(bin) + sin(bout) ] = l
One spot hopelessly lost,
but two spots get through
 interference  image formed!
Resolution (smallest resolvable d)
with incoherent illumination (all possible illumination directions):
dmin = l/(NAobj + NAcondenser )
l/2 NA if NAcondenser  NAobj (“Filling the back focal plane”)
bin
Two spots get through if
bout < a and bin < a.
d

## Slide 48
Aperture, Resolution & Contrast
Can adjust the condenser NA with the aperture iris
Sample
Objective
Tube lens
Imaging
path
Aperture iris
Field iris
Light source
Illumination
path
Collector
Condenser lens
Field lens
Back aperture
Intermed. image
Q: Don’t we always want
it full open??
A: No
Why? Tradeoff:
resolution vs. contrast

## Slide 49: NA and Resolution
High NA Objective Low NA Objective

## Slide 50
Alternate Definitions of Resolution
As the Full Width at Half Max
(FWHM) of the PSF
As the diameter of the Airy disk
(first dark ring of the PSF)
= “Rayleigh criterion”
(Probably most common definition)
Airy disk radius
≈ 0.61  /NA
FWHM
≈ 0.353  /NA

## Slide 51: Objective Types
Field flatness
Plan or not
Phase rings for phase contrast
Positive or negative
Diameter of ring (number)
Special Properties
Strain free for Polarization or DIC
Features
Correction collar for spherical aberration
Iris
Spring-loaded front end
Lockable front end
Basic properties
Magnification
Numerical Aperture (NA)
Infinite or finite conjugate
Cover slip thickness if any
Immersion fluid if any
Correction class
Achromat
Fluor
Apochromat

## Slide 52: Further reading
www.microscopyu.com
micro.magnet.fsu.edu
Douglas B. Murphy “Fundamentals of Light Microscopy and Electronic Imaging”
James Pawley, Ed. “Handbook of Biological Confocal Microscopy, 3rd ed.”
Acknowledgements
Ron Vale / Mats Gustafsson