## Slide 1: EMCCD Cameras
Rachit Mohindra
Photometrics

## Slide 2: EMCCDs – The Need for Low Light Imaging
Low excitation energy to reduce photo-bleaching/toxicity
  Live cell fluorescence
  Fast bleaching dyes
Samples in ultra-low concentrations
  ~10-10 M in typical single molecule fluorescence applications
Observe Processes with fast kinetics
  Motility studies
  Ion ratio imaging
Essentially - Low Light and Speed

## Slide 3: CCD Image Acquisition
Frame Transfer CCD

## Slide 4: EMCCD Cameras
Normal Voltage Serial Register
High Voltage Serial Register
Back-illuminated, frame-transfer CCD (for >90% QE) or standard front-illuminated (65% QE)

## Slide 5: EMCCD Cameras

## Slide 6: How Does this Help?
If you are trying to measure signal that is lower than the read noise – You’re not going to see anything
If you multiply up the signal over the read-noise – Then you’ll see something.
But – the signal you’re measuring isn’t actually changing.
You’re effectively reducing the read noise

## Slide 7: Reduction of Read Noise
Apply EM Gain factor of 200X
Boost Signal 200X
Equivalent to reducing the read noise by a factor of 200X
If Read Noise is 50e-, after EM gain, it is equal to 0.25e-
Allows extremely low-light detection

## Slide 8: Detect Signals Below the Read-Noise
Without EM Gain
With EM Gain

## Slide 9: Downfalls of EMCCDs
Clock Induced Charge increases
  Time independent (unlike dark current)
  Increases background events
Excess Noise Factor
  Increases noise by a factor of ~1.4
  “Halves the Quantum Efficiency” – Not exactly
Loss of Full Well Capacity
  Inversely proportional to the amount of EM Gain applied
Aging of EM Gain

## Slide 10: Excess Noise Factor
Increases only Shot Noise and Dark Noise by a factor of ~√2 or ~1.4
Important at higher light levels, where images are dominated by Shot Noise

## Slide 11: Calibrate your EM Gain
As EM Gain ages – setting 300X doesn’t always provide 300X…
Most Cameras provide you the ability to calibrate

## Slide 12: The Sweet Spot
Maximum EM Gain is not the best setting
  Maximizes Excess Noise
  Maximizes Clock Induced Charge/Spurious noise
  Drastically reduced Dynamic Range
  Read Noise asymptotically reaches a minimum Read Noise
Reduce the Read Noise to <1e-
  No more EM Gain needed
  Maximize the sensitivity, minimize the noise
  Extend the life of the EM Gain

## Slide 13
With increasing EM Gain – ADU’s increase, But Signal detected by the sensor remains constant
After a certain point, the electron count will remain the same
Chart: EM Gain
| 1 | 50 | 100 | 200 | 400 | 800
ADU corrected | 45 | 1539 | 2978 | 6393 | 13770 | 24708
Chart: EM Gain
| 1 | 50 | 100 | 200 | 400 | 800
electrons | 275 | 182 | 153 | 155 | 162 | 157

## Slide 14: EMCCD’s – A Summary
Low Light Imaging that requires Speed
Effectively reduces your EM Gain
Not a perfect technology
Find the Sweet Spot – Don’t need more EM than this

## Slide 15
Thank You!
For more information please visit:-
  www.photometrics.com