## Slide 1
(Image: T. Wittman, Scripps)
Principles & Practice of
Light Microscopy: 2

## Slide 2: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Diffraction spot
on image plane
= Point Spread Function
Sample
Objective

## Slide 3: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Sample
Objective
Diffraction spot
on image plane
= Point Spread Function

## Slide 4: Aperture and Resolution
Tube lens
Back focal plane aperture
Intermediate image plane
Sample
Objective
Diffraction spot
on image plane
= Point Spread Function

## Slide 5: Aperture and Resolution
Image resolution improves with Numerical Aperture (NA)
Sample
Objective
Tube lens
Back focal plane aperture
Intermediate image plane

NA = n sin()
 = light gathering angle
n = refractive index of sample
where:
Diffraction spot
on image plane
= Point Spread Function

## Slide 6
aobjective
For the highest resolution,
we need to have
condenser  objective
NAcondenser  NAobjective
with oil immersion objectives,
we need an oil immersion condenser!
acondenser
Filling the back focal plane
In trans-illumination microscopy, to get maximum resolution,
the illumination must “fill the back focal plane”
Objective
Condenser
Light source
Back
focal
plane

## Slide 7
Aperture, Resolution & Contrast
Can adjust the condenser NA with the aperture iris
Sample
Objective
Tube lens
Imaging
path
Aperture iris
Field iris
Light source
Illumination
path
Collector
Condenser lens
Field lens
Back aperture
Intermed. image
Q: Don’t we always want
it full open??
A: No
Why? Tradeoff:
resolution vs. contrast

## Slide 8
Spatial frequencies & the
Optical Transfer Function (OTF)
Object
Observed
image
(Spatial frequency,
periods/meter)
k
OTF(k)
1
(Image
contrast)
Resolution limit:
kmax = 2 NA / l

## Slide 9
Resolution & Contrast vs. Illumination aperture
Resolution
Contrast
NAcondenser  0
NAcondenser  NAobj
Increasing the illumination aperture
increases resolution
but decreases contrast
(“Coherent
illumination”)
(= Full aperture,
“incoherent
illumination”)
Pupil
appearance
Resolution limit:
kmax = (NAobjective+ NAcondenser)/ l

## Slide 10
Definitions of Resolution
|k|
OTF(k)
1
Cutoff frequency
kmax = 2 NA / l
As the OTF cutoff frequency
As the Full Width at Half Max
(FWHM) of the PSF
As the diameter of the Airy disk
(first dark ring of the PSF)
= “Rayleigh criterion”
Airy disk diameter
≈ 0.61  /NA
FWHM
≈ 0.353  /NA
1/kmax
= 0.5  /NA

## Slide 11
Measured
x
z
Calculated
The 3D Point Spread Function (PSF)
Z=0
Z=+2µm
Z=-2µm
2D PSF
for different defocus
The image of a point object
x
y
3D PSF

## Slide 12
x - z

## Slide 13: Z-resolution, a.k.a. depth of field, for widefield microscopy
NA | Resolution (nm; X-Y) | depth of field (mm)
0.3 | 1017 | 11.1
0.75 | 407 | 1.77
0.95 | 321 | 1.11
1.4 | 218 | 0.773
Z-resolution:
2ln / NA2

## Slide 14: Summary: Numerical Aperture and Resolution
Numerical aperture, not magnification, sets the smallest details you can resolve in an image
Increasing NA also increases the amount of light collected by the lens, thereby increasing the brightness of the image – this scales as NA2

## Slide 15: Specifications for some common objectives
Magnification | NA | Resolution (nm) | Depth of Field (nm) | Light gathering (arb. units)
10 | 0.3 | 1017 | 16830 | 0.09
20 | 0.75 | 407 | 2690 | 0.56
40 | 0.95 | 321 | 1680 | 0.90
40 | 1.3 | 235 | 896 | 1.69
60 | 1.2 | 254 | 926 | 1.44
60 | 1.4 | 218 | 773 | 1.96
100 | 1.4 | 218 | 773 | 1.96

## Slide 16
Aberrations
They are the enemy

## Slide 17
Aberrations
Chromatic aberrations
  Longitudinal chr. Ab.
  Lateral chr. Ab.
Wavefront aberrations
  Spherical aberration
  Astigmatism
  Coma
  …
Curvature of field
Distortion

## Slide 18: Wavefront Aberrations
Aberrated wavefront
in the pupil
Ideal wavefront
in the pupil

## Slide 19
Wavefront Aberrations
(piston)
(tilt)
Astigmatism
Defocus
Coma
Trefoil
Spherical ab.
Secondary coma
Secondary
spherical ab.

## Slide 20
(piston)
(tilt)
Astigmatism
Defocus
Coma
Trefoil
Spherical ab.
Secondary coma
Secondary
spherical ab.
PSF Aberrations

## Slide 21: Spherical Aberration

## Slide 22
Spherical Aberration
Point spread functions
z
x
x
Ideal
1 wave of spherical ab

## Slide 23: Causes of spherical aberration
Modern objectives are complicated!
The optical design requires specifying the optical path length between the sample and the back focal plane of the lens
OPL = l1n1 + l2n2 + … + lnnn

## Slide 24
Design compromises
Manufacturing tolerances
Immersion fluid index error
Temperature variation
Cover slip thickness
  (high-NA objectives except oil immersion)
Correction collar setting
Sample refractive index mismatch
Sources of Spherical Aberration

## Slide 25
Index Mismatch
& Spherical Aberration
objective
Immersion
fluid
n1
Cover glass
n2
Sample
Spherical aberration
unless n2 = n1
Focus
into sample
Focus at cover slip

## Slide 26
Index Mismatch
& Spherical Aberration
n1=1.515 (oil)
n2=1.44
(Vectashield)
z=0 µm
z=25 µm
z=50 µm

## Slide 27: How to recognize spherical aberration
Unaberrated
Aberrated
0.5 mm steps
1 mm steps

## Slide 28: What can you do about spherical aberration?
Use 0.17 mm coverslips (~ #1.5)
Work close to the coverslip
Match lenses to the refractive index of your samples, and vice versa
  For aqueous samples, use water immersion / water dipping lenses
  For fixed samples and oil immersion lenses, mount your sample in a medium with n = 1.515
Adjust objective correction collar when available
Use lower NA lenses

## Slide 29: Correction collars

## Slide 30: Correction collars
A correction collar can only eliminate spherical aberration at a single focal plane

## Slide 31: Example
Aberrated
Unaberrated

## Slide 32
n1
Mechanical
focus
step zm
n2
Optical
focus
step zo
Index Mismatch & Axial Scaling
If there is index mismatch,
your z pixel size is not what you think

## Slide 33: Index Mismatch & Axial Scaling
2 m beads
n = 1.32
xy
z
xy
z
n = 1.52
Mark Chan

## Slide 34
Off-axis (edges of field of view)
Sources of Astigmatism & Coma
On-axis (center of field of view)
All objectives have some
Present in the design
You get what you pay for
Should be none, by symmetry.
If they are there, they could be from:
  manufacturing or assembly tolerances
  dirt or abuse
  Misalignment (tilt, off-axis shift of something)
  bad downstream components (mirrors, dichroics, filters…)
  Air bubble in the immersion fluid or sample
  Tilted cover slip
    (dry and water-immersion high-NA lenses)

## Slide 35
More about
Spatial frequencies & the
Optical Transfer Function (OTF)

## Slide 36
The response to pure waves is well-defined by the
Optical Transfer Function (OTF)
Object
Observed
image
(Spatial frequency,
periods/meter)
k
OTF(k)
1
(Image
contrast)
Resolution limit:
kmax = 2 NA / l

## Slide 37
Think of Images as Sums of Waves
another wave
one wave
(2 waves)
+
=
(10000 waves)
+ (…) =
… or “spatial frequency components”
(25 waves)
+ (…) =

## Slide 38
Frequency Space
Frequency (how many periods/meter?)
Direction
Amplitude (how strong is it?)
Phase (where are the peaks & troughs?)
period
direction
To describe a wave,
we need to specify its:
ky
kx
Distance from origin
Direction from origin
Magnitude of value
Phase of value
Can describe it by
a value at a point
complex
k = (kx , ky)

## Slide 39
ky
kx
Frequency Space
ky
kx
and the Fourier Transform
Fourier
Transform

## Slide 40
Properties of the Fourier Transform
Symmetry:
The Fourier Transform of the Fourier Transform
is the original image
Completeness:
The Fourier Transform contains all the information
of the original image
Fourier
transform

## Slide 41
The OTF and Imaging
Fourier
Transform
True
Object
Observed
Image
OTF

=
=
?
?

convolution
PSF

## Slide 42
Convolutions
(f  g)(r) = f(a) g(r-a) da
Why do we care?
They are everywhere…
The convolution theorem:
If
then
h(r) = (fg)(r),
h(k) = f(k) g(k)
A convolution in real space becomes
a product in frequency space & vice versa
So what is a convolution, intuitively?
“Blurring”
“Drag and stamp”
=
f
g
fg
=
x
x
y
x
y
y
Symmetry: g  f = f  g



## Slide 43
Observable
Region
ky
kx
The Transfer Function Lives in Frequency Space
Object
|k|
OTF(k)
Observed
image

## Slide 44
The 2D In-focus Optical Transfer Function (OTF)
|k|
OTF(k)
ky
kx
OTF(k)
(Idealized calculations)

## Slide 45
?
?
2D PSF
3D PSF
2D OTF
The 3D OTF
2D F.T.
3D OTF
3D F.T.

## Slide 46
Values of the 3D OTF
kz
kx

## Slide 47
3D Observable Region
kz
ky
kz
ky
= OTF support
= Region where the OTF is non-zero

## Slide 48
kz
kx
Kxymax =
2 n sin(a) / l
= 2 NA / l
So what is the resolution?
“Missing
Cone” of
information
Kzmax =
n (1-cos(a)) / l

## Slide 49
kz
kx
Kzmax =
n (1-cos(a)) / l
Kxymax =
2 n sin(a) / l
= 2 NA / l
So what is the resolution?
“Missing
Cone” of
information
Lowering the NA
Degrades the
axial resolution
faster than the
lateral resolution
But low axial resolution
= long depth of field
This is good,
if 2D is enough

## Slide 50
kz
kx
Kzmax =
n (1-cos(a)) / l
Kxymax =
2 n sin(a) / l
= 2 NA / l
NA = 1.4
n=1.515
a = 67.5°
l = 600 nm
Lateral (XY) resolution:
1/ Kxymax = 0.21 µm
Axial (Z) resolution:
1/ Kzmax = 0.64 µm
So what is the resolution?
Example:
a high-end objective
“Missing
Cone” of
information

## Slide 51
Nomenclature
Optical Transfer Function, OTF
  Complex value with amplitude and phase
Contrast Transfer Function, CTF
Modulation Transfer Function, MTF
  Same thing without the phase information

## Slide 52: Resources
Slides available at: http://nic.ucsf.edu/edu.html
http://www.microscopyu.com
http://micro.magnet.fsu.edu
Douglas B. Murphy “Fundamentals of Light Microscopy and Electronic Imaging”
James Pawley, Ed. “Handbook of Biological Confocal Microscopy, 3rd ed.”
Acknowledgements
Steve Ross, Mats Gustafsson